## Slide 1
INTRODUCTION TO MRI
May 8, 2018
Structural MRI
Lynne J. Williams, Ph.D.
BCCH MRI Research Facility

## Slide 2
LAB SESSIONS
Wednesdays (11 to 12)
https://bluejeans.com/320270335
Fridays (10 to 11)
https://bluejeans.com/663641726

## Slide 3
REVIEW
Radio frequency (B1) waves knock over the aligned, precessing protons. When the protons realign with B0, they release energy at the resonance (Larmor) frequency
PROTON RELEASE ENERGY
Spin (aka angular momentum) generates a magnetic moment that allows protons to precess at a specific resonance (Larmor) frequency in an external magnetic field (B0)
PROTONS PRECESS
The relaxation rates (return to alignment with B0) are recorded in K-space as a spatial frequencies (number of cycles per meter)
DATA GETS STORED IN K-SPACE

## Slide 4
REVIEW
QUESTIONS?

## Slide 5
WEEK 2
STRUCTURAL MRI
TODAY’S CLASS
What is relaxation (and why we need to know)?
What do MR parameters mean?
Why should we care about Quality Control?
MRI Scans gone bad (aka Artifacts)

## Slide 6
WEEK 2
STRUCTURAL MRI
RELAXATION
T1, T2 & T2*
STRUCTURAL MRI
Lynne J. Williams, Ph.D.
BCCH MRI Research Facility

## Slide 7
T1 (LONGITUDINAL) VS
T2 (TRANSVERSE)
RELAXATION
T1 and T2 are magnetic timing parameters that differ between tissue types
  They are used as a source of contrast
The T1 and T2 time constants dictate the shape of the exponential recovery and decay curves of the longitudinal and transverse magnetization
Slide / 01
www.yourwebsite.com
T1
T2

## Slide 8
RECALL
RF PULSES
“KNOCK OVER”
PROTONS
Protons lined up with B0 get knocked over by an RF pulse
M is the net magnetization of all protons
M0 is the maximum net magnetization
B0

## Slide 9
63% recovery
T1 time constant

## Slide 10
37% decay
T2 time
constant

## Slide 11
The decay of Mxy and the regrowth of Mz should be considered separate processes.
The net magnetization vector is the vector sum of these two processes.
Mxy decay occurs faster than Mz recovery
Mxy decay can occur without Mz recovery, but the opposite is not true
DECAY AND
REGROWTH ARE
SEPARATE
PROCESSES
Slide / 01
www.yourwebsite.com

## Slide 12
T2 and T2*
T2
T2*
T2 dephasing includes only
irreversible sources of loss
of phase coherence
T2* dephasing also includes additional sources of loss of
phase coherence (e.g.,
magnetic field
inhomogeneities and
susceptibility)
Notes: T2 dephasing
Spin-spin interactions at the molecular level changes the frequency of precession and protons move out of phase
The spins (protons) go out of phase in the transverse plane when the Rf pulse is switched off. This dephasing is due to static magnetic field inhomogeneities intrinsic (i.e. susceptibility changes at tissue interfaces) and extrinsic (i.e. field inhomogeneities of the main magnetic field) to the examined object.
T2*
This causes changes to the frequency of precession at the molecular level and protons move out of phase even quicker

## Slide 13
(Free Induction Decay)
THEREFORE, FID
OSCILLATES AT THE
RESONANCE (LARMOR)
FREQUENCY, BUT IS
DAMPENED
BY T2* DECAY
ADDITIONAL DEPHASING OCCURS DUE TO MAGNETIC FIELD INHOMOGENEIT-IES AND SUSCEPTIBILITY

## Slide 14
WEEK 2
STRUCTURAL MRI
SCAN PARAMETERS
STRUCTURAL MRI

## Slide 15
MRI
PARAMETERS

## Slide 16
MRI
PARAMETERS
Slide / 01
Amount of rotation of the magnetization vector produced by an RF pulse (B1) with respect to the direction of the static magnetic field (B0).
FLIP ANGLE
- Hendrick RE (Ed). Glossary of MR Terms (5th Ed). Reston, VA: American College of Radiology 2005

## Slide 17
PARTIAL
VS 90°
FLIP
ANGLE
Image from Revise MRI

## Slide 18
MRI
PARAMETERS
An echo is the emission of energy in the form of an electromagnetic resonance signal of nuclei after its excitation.
ECHO
MR Technology Information Portal
Notes: Partial echo vs full echo. Partial echo is when the pre-dephasing lobe of the readout gradient is reduced, meaning the readout won't start as far out in k-space. Then the missing data points are synthesised using the conjugate symmetry of k-space. The result is decreased SNR, but a shorter minimum TE because less time is used in the readout before k0 is acquired.

## Slide 19
SPIN
ECHO
The spin echo pulse sequence has a 90° excitation pulse and refocuses some of the dephasing with a 180° refocusing RF pulse.
Image from Revise MRI
Notes: The spin echo pulse sequence has a 90° excitation pulse (GE has a small excitation pulse), and SE refocuses some of the dephasing which occurs during the echo time using a 180° refocusing RF pulse. A GE does not have the 180°.

## Slide 20
GRADIENT SPIN
ECHO
Image from Revise MRI

## Slide 21
MRI
PARAMETERS
Slide / 01
The echo time (TE) represents the time from the centre of the RF-pulse to the centre of the echo.
- Revise MRI
ECHO TIME (TE)

## Slide 22
REPETITION TIME (TR)
& ECHO TIME (TE)

## Slide 23
In inversion recovery, time between middle of
inverting (180°) RF pulse and middle of the subsequent exciting (90°) pulse to detect amount of longitudinal magnetization
MRI
PARAMETERS
Slide / 01
INVERSION TIME (TI)
- Hendrick RE (Ed). Glossary of MR Terms (5th Ed). Reston, VA: American College of Radiology 2005

## Slide 24
Notes: The usual Inversion Recovery sequence begins with a 180°-inverting pulse
  Reverses the longitudinal magnetization (Mz) for all tissues.
During the inversion time (TI) interval
  Mz regrows via T1-relaxation to restore their equilibrium alignment in the +z-direction.
A spin echo (or more commonly a fast spin echo) sequence then generates a signal based on the longitudinal magnetization of each tissue at time TI
For tissues with Mz close to zero at time TI
  Tissue produces little or no signal
  Makes no significant contribution to the magnitude-reconstructed image

## Slide 25
MRI
PARAMETERS
Slide / 01
The rectangular region superimposed over the human body over which MRI data are acquired. Its dimensions are specified in length in each in-plane direction and are controlled by the application of frequency encode and phase-encode gradients.
FIELD OF VIEW (FOV)
- Hendrick RE (Ed). Glossary of MR Terms (5th Ed). Reston, VA: American College of Radiology 2005

## Slide 27
$ fslinfo STRUCT0000.nii.gz
data_type INT16
dim1 256
dim2 256
dim3 256
dim4 1
datatype 4
pixdim1 0.900000
pixdim2 0.900000
pixdim3 0.900000
pixdim4 1.000000
cal_max 928.0000
cal_min 0.0000
file_type NIFTI-1+
MATRIX SIZE
x = FOVx/Nx
y = FOVy/Ny
z3D = FOVz/Nz
z2D = slice
thickness
VOXEL DIMENSIONS

## Slide 28
MRI
PARAMETERS
Slide / 01
The thickness of a slice. As the slice profile may not be sharp edged, a criterion such as the distance between the points at half the sensitivity of the maximum (FWHM) or the equivalent rectangular width may be useful.
SLICE THICKNESS
- Hendrick RE (Ed). Glossary of MR Terms (5th Ed). Reston, VA: American College of Radiology 2005

## Slide 29
1 mm × 1 mm × 1 mm
.97 mm ×.97 mm × 5 mm
ISOTROPIC
ANISOTROPIC
X
Y
Z
X
Y
Z

## Slide 30
1 mm × 1mm × 1mm
$ fslinfo STRUCT0005.nii.gz
data_type INT16
dim1 174
dim2 192
dim3 224
dim4 1
datatype 4
pixdim1 1.000000
pixdim2 1.000000
pixdim3 1.000000
pixdim4 2.100000
cal_max 1931.0000
cal_min 0.0000
file_type NIFTI-1+

## Slide 31
.97 mm ×.97 mm × 5 mm
$ fslinfo STRUCT0007.nii.gz
data_type INT16
dim1 256
dim2 256
dim3 30
dim4 1
datatype 4
pixdim1 0.976562
pixdim2 0.976562
pixdim3 5.000000
pixdim4 1.000000
cal_max 4095.0000
cal_min 0.0000
file_type NIFTI-1+

## Slide 32
WEEK 2
STRUCTURAL MRI
QUALITY CONTROL
STRUCTURAL MRI
Lynne J. Williams, Ph.D.
BCCH MRI Research Facility

## Slide 33
RULE #1:
A GOOD OFFENCE
MAKES A
GOOD DEFENCE
CONTROL FOR
MOTION AT
THE SOURCE

https://hub.bcchr.ca/display/equip/MRI+Research+Facility+Frequently+Asked+Questions

## Slide 34
MOTION
ARTIFACT

## Slide 35
RULE #2:
LOOK AT YOUR
DATA
THIS CAN DETERMINE STEPS IN YOUR PROCESSING PIPELINE


## Slide 36
BIAS

## Slide 37
GRADIENT
NONLINEARITY
ARTIFACT
Expansion in the inferior regions

## Slide 38
RULE #3:
LOOK AT YOUR
DATA
DETERMINES WHETHER THE DATA SHOULD BE INCLUDED IN THE ANALYSIS


## Slide 39
WRAP
ARTIFACT
Some of the signal from the neck overlaps with portions of the brain

## Slide 40
RF ZIPPER
ARTIFACT

## Slide 41
LOW
SNR

## Slide 42
RECON
ARTIFACT
Poor SNR in central regions due to high acceleration

## Slide 43
(ANOTHER)
RECON
ARTIFACT
High acceleration causing reconstruction artifacts

## Slide 44
Magnetic Susceptibility Artifacts on MRI: A Hairy Situation
Robert C. McKinstry, III and Delma Y. Jarrett
American Journal of Roentgenology 2004 182:2, 532-532

## Slide 45
WEEK 2
STRUCTURAL MRI
NAME THE ARTIFACT
STRUCTURAL MRI
Lynne J. Williams, Ph.D.
BCCH MRI Research Facility
Artifact images from http://www.neuroimagingprimers.org/

## Slide 46
MOTION

## Slide 47
WRAP

## Slide 48
MOTION

## Slide 49
MOTION

## Slide 50
BIAS
CAN BE
CORRECTED
DURING
PREPROCESSING

## Slide 51
NONE
ENLARGED
VENTRICLES
Notes: None, enlarged ventricles

## Slide 52
NONE

## Slide 53
BIAS
CAN BE
CORRECTED
DURING
PREPROCESSING

## Slide 54
RF
ZIPPER

## Slide 55
REVIEW
Radio frequency (B1) waves knock over the aligned, precessing protons. When the protons realign with B0, they release energy at the resonance (Larmor) frequency
PROTON RELEASE ENERGY
Spin (aka angular momentum) generates a magnetic moment that allows protons to precess at a specific resonance (Larmor) frequency in an external magnetic field (B0)
PROTONS PRECESS
The relaxation rates (return to alignment with B0) are recorded in K-space as a spatial frequencies (number of cycles per meter)
DATA GETS STORED IN K-SPACE

## Slide 56
REVIEW
@lukenxie

## Slide 57
REVIEW
T2
T1
@lukenxie

## Slide 58
REVIEW

## Slide 59
REVIEW
QUALITY
CONTROL

## Slide 60
REVIEW
QUESTIONS?

## Slide 61
WEEK 2
STRUCTURAL MRI
BACK TO
QUALITY CONTROL
STRUCTURAL MRI
Lynne J. Williams, Ph.D.
BCCH MRI Research Facility

## Slide 62
RULE #4:
LOOK AT YOUR
DATA
QUALITATIVE AND QUANTITATIVE MEASURES NEEDED


## Slide 63
MRIQC’S PROCESSING DATA FLOW
Esteban O, Birman D, Schaer M, Koyejo OO, Poldrack RA, et al. (2017) MRIQC: Advancing the automatic prediction of image quality in MRI from unseen sites. PLOS ONE 12(9): e0184661. https://doi.org/10.1371/journal.pone.0184661
http://journals.plos.org/plosone/article?id=10.1371/journal.pone.0184661

## Slide 64
VISUAL REPORTS
Esteban O, Birman D, Schaer M, Koyejo OO, Poldrack RA, et al. (2017) MRIQC: Advancing the automatic prediction of image quality in MRI from unseen sites. PLOS ONE 12(9): e0184661. https://doi.org/10.1371/journal.pone.0184661
http://journals.plos.org/plosone/article?id=10.1371/journal.pone.0184661

## Slide 65
Notes: Signal-to-Noise Ratio (SNR) [snr]: The mean intensity within gray matter divided by the standard deviation of the values outside the brain. Higher values are better 1.
Contrast to Noise Ratio (CNR) [cnr]: The mean of the gray matter intensity values minus the mean of the white matter intensity values divided by the standard deviation of the values outside the brain. Higher values are better 1.
Foreground to Background Energy Ratio [fber]: The variance of voxels inside the brain divided by the variance of voxels outside the brain. Higher values are better.
Percent Artifact Voxels (Qi1) [qi1]: The proportion of voxels outside the brain with artifacts to the total number of voxels outside the brain. Lower values are better 2.
Smoothness of Voxels (FWHM) [fwhm, fwhm_x, fwhm_y, fwhm_z]: The full-width half maximum (FWHM) of the spatial distribution of the image intensity values in voxel units. Lower values are better 3.
Entropy Focus Criterion (EFC) [efc]: The Shannon entropy of voxel intensities proportional to the maximum possibly entropy for a similarly sized image. Indicates ghosting and head motion-induced blurring. Lower values are better

## Slide 66
WEEK 2
STRUCTURAL MRI
PROCESSING PIPELINES
STRUCTURAL MRI
Lynne J. Williams, Ph.D.
BCCH MRI Research Facility

## Slide 68
HCP PRE-FREESURFER PIPELINE

## Slide 69
HCP FREESURFER PIPELINE

## Slide 70
© 2018 BC CHILDREN’S HOSPITAL MRI RESEARCH FACILITY
This work is licensed under the Creative Commons Attribution-NONCOMMERCIAL-ShareAlike 4.0 International License. To view a copy of this license, visit http://creativecommons.org/licenses/by-nc-sa/4.0/..
THANK YOU
BCCH MRI Research Facility

## Slide 71
NONE
ENLARGED
VENTRICLES
Notes: None, enlarged ventricles

## Slide 72
NONE
ENLARGED
VENTRICLES
Notes: None, enlarged ventricles

## Slide 73
PROTON DENSITY
T1 provides insight into the repetition time
T2 is directly related to the lifetime of the excited state

## Slide 74
T1 vs T2
T1 provides insight into the repetition time
T2 is directly related to the lifetime of the excited state

## Slide 75
Slide / 01
www.yourwebsite.com
Note that the T2 FLAIR scan came from a different scan session (although all from the same individual)
EXAMPLES
OF COMMON
STRUCTURAL
SCANS
T1
T2
PD
T2 FLAIR

## Slide 76
YOUR TITLE
HERE
Far far away, behind the word mountains, far from the countries Vokalia and Consonantia, there live the blind texts. Separated they live in Bookmarksgrove right at the coast of the Semantics, a large language ocean.

## Slide 77
Far far away, behind the word mountains, far from the countries Vokalia and Consonantia, there live the blind texts. Separated they live in Bookmarksgrove right at the coast of the Semantics, a large language ocean.
- The Quotes
Far far away, behind the word mountains, far from the countries Vokalia and Consonantia, there live the blind texts. Separated they live in Bookmarksgrove right at the coast of the Semantics, a large language ocean.
- The Quotes
YOUR
TITLE HERE
Notes: TR and TE HERE!!!!!

## Slide 78
Sign
T1
T2
PD
Short TE
Long TE
Short TR
Long TR
Poor CNR & SNR
Not done in practice
EXAMPLES
OF HOW SCAN
MODALIES
RELATE TO
TR & TE

## Slide 79
Slide / 01
ACCEPTABLE
T1 SCANS
Far far away, behind the word mountains, far from the countries Vokalia and Consonantia, there live the blind texts.
Chart: 
| 2014 | 2015 | 2016 | 2017
Series 1 | 4.3 | 2.5 | 3.5 | 4.5
Series 2 | 2.4 | 4.4 | 1.8 | 1
Series 3 | 2 | 2 | 3 | 5
series 4 | 1 | 3 | 2.7 | 3.7
YOUR TITLE
85%
Far far away, behind the word mountains, far from the countries Vokalia and Consonantia, there live the blind texts.

## Slide 80
© 2018 BC CHILDREN’S HOSPITAL MRI RESEARCH FACILITY
This work is licensed under the Creative Commons Attribution-NONCOMMERCIAL-ShareAlike 4.0 International License. To view a copy of this license, visit http://creativecommons.org/licenses/by-nc-sa/4.0/..
THANK YOU
BCCH MRI Research Facility